# User manual for A behavioural setup to study temperature response of ciona intestinalis larvae

## Before testing
1 . Before Testing
1. Clean the chamber
   Remove any dust or residue from the habitat with filtered seawater and a lint-free wipe. Let it air-dry.
2. Verify the temperature probes
   Ensure both probes are seated firmly against the chamber walls and their cables are undamaged.
3. Power the controller
   Connect the USB cable between the computer and the Arduino Uno. The controller screen will light up.
4. Set the target temperature
   Use the on-screen menu to dial in the desired temperature before energising the Peltier elements.
5. Switch on the system
   Flip the main ON/OFF switch. The LEDs illuminate, and the Peltiers begin driving toward the set point.
6. Confirm a sensible starting value
   Check that the temperature shown on the display is realistic (i.e., within a few degrees of ambient). This prevents the Peltiers from “running away” toward an erroneous set point.
7. Seat the habitat correctly
   Place the habitat so it rests flat and makes full contact with both Peltier faces.

## During testing
1. Temperature ramping
   Large jumps take time. When raising the bath from a low (e.g., 10 °C) to a high (e.g., 40 °C) temperature, increase in small steps—pause at an intermediate value to avoid overshoot. For changes < 20 °C, a direct adjustment is usually safe.
2. Stability checks
   Check if the temperature change are working and that no problems occur.
3. Heat-sink inspection
   For multi-hour runs, feel the heat-sinks. They should be warm, not hot. If they are hot to the touch, verify that the cooling fans are spinning freely and unobstructed.

## After testing

Power down
After testing turn of the powersupply and disconnect the USB from the computer.

Empty and wash
Remove the contents of the habitat and wash it out.